# VIS-NIR Spectrometer Data Processing Pipeline

Presentation-ready review and slide wording

Aligned to the current PythonGUI and STM32 USB CDC frame stream on 2026-04-29.

Purpose: define the real data path from CCD readout through display/export, and provide short copy that can be placed directly on the slide deck.

# Recommended Slide Revision

The provided five-phase slide is directionally correct, but several labels should be updated so the presentation matches the current implementation.

| Current slide wording | Recommended wording | Reason |
| --- | --- | --- |
| Microcontroller Readout | Microcontroller Readout | Keep the title, but make the body more specific: STM32 clocks the CCD, captures a 3694-sample ADC line with DMA, and sends a CCD1 binary USB CDC frame. |
| Dark Frame Subtraction | Bias and Dark Correction | The current code supports master bias B_p, a frame-wise dark reference from shielded pixels 16-28, and optional master-dark offsets. It is broader than a simple dark-frame subtraction. |
| Wavelength Calibrations | Wavelength Calibration | Use singular wording. The implemented workflow fits one pixel-to-wavelength polynomial from reference mapping points. |
| QE and Intensity Normalization | Response Correction and Normalization | This phase includes flat-field/PRNU correction, QE/response correction, and final display normalization. |
| Spectrum Display | Display, Spectrogram, and Export | The GUI now shows either a line spectrum or rolling spectrogram and exports CSV with raw plus derived columns. |

Notation note for slides: use "lambda = wavelength; a.u. = arbitrary units" or keep the Greek lambda symbol if the slide font supports it cleanly.

# Pipeline At A Glance

| Phase | Slide title | Short slide statement |
| --- | --- | --- |
| 1 | Microcontroller Readout | STM32 clocks the CCD, captures a 3694-sample ADC line, and streams it as a CCD1 USB packet. |
| 2 | Bias and Dark Correction | Bias and dark baselines are removed so the remaining signal tracks light intensity. |
| 3 | Wavelength Calibration | Known spectral lines map pixel index to wavelength with a fitted polynomial. |
| 4 | Response Correction and Normalization | Flat-field and QE corrections compensate instrument response, then the signal is scaled for display. |
| 5 | Display, Spectrogram, and Export | The GUI renders the calibrated spectrum or spectrogram and exports raw plus processed data. |

# Slide Copy: One-Line Version

Use these statements directly on the overview slide cards.

- 1. Microcontroller Readout - CCD signal captured by STM32 ADC + DMA and sent as a binary USB frame.

- 2. Bias and Dark Correction - Bias and dark baselines are removed from each pixel signal.

- 3. Wavelength Calibration - Reference peaks convert pixel position into wavelength.

- 4. Response Correction and Normalization - Flat-field and QE corrections are applied before scaling.

- 5. Display, Spectrogram, and Export - Python displays live spectra and saves raw plus processed data.

# Phase 1 - Microcontroller Readout

Slide statement: CCD signal captured by STM32 ADC + DMA and sent as a binary USB frame.

Presenter note: The STM32 firmware generates the CCD timing, captures the analog output with ADC1 using DMA, and streams each frame to the desktop app as a structured CCD1 packet over USB CDC.

| Technical definition | Slide/design guidance |
| --- | --- |
| - TIM1 provides the PHI/M CCD clock, TIM2 triggers ADC conversions, and TIM4 defines the ICG-synchronized frame boundary. - Each normal packet contains 3694 ADC samples: 32 leading dummy samples, 3648 effective pixels, and 14 trailing dummy samples. - The USB packet includes frame_id, sample_count, effective_start, effective_count, flags, and the little-endian 16-bit ADC payload. - The Python app preserves the raw ADC samples before applying calibration for display and export. | - Replace "USB serial stream" with "USB CDC CCD1 binary frame" for precision. - In the right-hand schematic placeholder, show CCD output -> ADC1 + DMA -> CCD1 packet -> Python parser. - Keep the three sub-steps from the example slide: CCD readout, ADC data acquisition, output to desktop. |

# Phase 2 - Bias and Dark Correction

Slide statement: Bias and dark baselines are removed from each pixel signal.

Presenter note: The desktop processing removes additive offsets before any multiplicative correction. The current implementation handles a stored master bias, a frame-wise dark reference from shielded pixels, and optional master-dark offsets.

| Technical definition | Slide/design guidance |
| --- | --- |
| - Raw ADC counts are converted to floating-point arrays for processing. - If stored bias counts exist, the app subtracts the master bias B_p from the raw frame. - When dark subtraction is enabled, the frame dark reference beta_f is the median of shielded pixels 16 through 28 after bias removal. - Because the current CCD readout is inverted, the light-tracking count is the distance from the dark baseline down to the sample, clipped at zero. - Optional master-dark offsets D_p can also be subtracted on a per-pixel basis. | - Rename the phase from "Dark Frame Subtraction" to "Bias and Dark Correction." - Use a before/after plot that shows baseline removal rather than implying a single saved dark frame only. - Short caption: "Remove B_p, frame dark beta_f, and optional D_p before response correction." |

# Phase 3 - Wavelength Calibration

Slide statement: Reference peaks convert pixel position into wavelength.

Presenter note: Known spectral reference peaks are captured as pixel,wavelength pairs. The calibration manager fits a polynomial so each sample index can be labeled in nanometers.

| Technical definition | Slide/design guidance |
| --- | --- |
| - The calibration manager stores pixel_mapping_points and wavelength_coefficients in the calibration JSON. - Guided mapping lets the user enter laser wavelengths, click the peak on the live plot, and capture the selected pixel. - The fit uses a configurable polynomial order, limited by the number of mapping points. - The x-axis remains in pixel mode until real mapping points and non-default coefficients have been saved. | - Change "Wavelength Calibrations" to "Wavelength Calibration." - The example labels 405 nm, 536 nm, and 650 nm are a good visual pattern if those match the actual reference sources used. - Show the x-axis changing from pixel index to lambda (nm). |

# Phase 4 - Response Correction and Normalization

Slide statement: Flat-field and QE corrections are applied before scaling.

Presenter note: After additive correction and wavelength assignment, the app applies multiplicative response corrections. This is where flat-field or PRNU terms and wavelength-dependent QE/response terms compensate the instrument response.

| Technical definition | Slide/design guidance |
| --- | --- |
| - Flat-field or PRNU correction multiplies the corrected signal by per-pixel factors when enabled. - QE/response correction builds an interpolated response curve from wavelength, response points and divides the signal by that curve. - The QE curve is normalized at the configured reference wavelength, currently defaulting to 550 nm. - Display normalization can be absolute saturation referenced or per-frame auto peak. - The rolling spectrogram uses a stable absolute-style normalization so colors remain comparable over time. | - Consider renaming this phase to "Response Correction and Normalization" so flat-field and QE both fit cleanly. - Show the corrected spectrum on a 0-to-1 normalized intensity axis. - Use "normalized intensity (a.u.)" if the y-axis has no physical radiometric unit yet. |

# Phase 5 - Display, Spectrogram, and Export

Slide statement: Python displays live spectra and saves raw plus processed data.

Presenter note: The Kivy desktop app renders the live processed spectrum, can switch to a rolling spectrogram, tracks session state and diagnostics, and exports CSV data for later analysis.

| Technical definition | Slide/design guidance |
| --- | --- |
| - For full frames, the GUI plots only the effective pixels so leading and trailing dummy samples do not dominate the view. - The live line plot uses processed display counts when available, while raw ADC counts remain stored in the frame. - The rolling spectrogram uses compact 256-column rows built from recent effective-pixel data. - CSV export includes raw_adc_count, processed_adc_count, frame_dark_reference_count, wavelength_nm, volts, and processed_intensity. - The Performance card can be enabled to inspect throughput, queue pressure, missed frames, and UI cadence. | - Rename the phase from "Spectrum Display" to "Display, Spectrogram, and Export" if there is room. - The mini GUI drawing should include live spectrum/spectrogram and CSV export, not only a run/stop control. - Short caption: "Render live data and preserve analysis-ready CSV output." |

# Speaker-Ready Pipeline Narrative

The VIS-NIR spectrometer pipeline begins in the STM32 firmware. The microcontroller generates the CCD timing, samples the analog output with ADC and DMA, and streams each captured line as a structured CCD1 USB CDC frame.

On the desktop side, the Python app reconstructs the incoming binary frames and preserves the raw ADC counts. Calibration starts by removing additive offsets: stored master bias, a frame-wise dark reference measured from shielded pixels, and optional master-dark offsets.

The corrected pixel signal is then assigned a wavelength using the saved pixel-to-wavelength calibration. Reference lines or laser peaks define mapping points, and the calibration manager fits a polynomial from those points.

After the wavelength map, the signal passes through multiplicative response correction. Flat-field or PRNU factors correct per-pixel sensitivity, and QE/response points correct wavelength-dependent throughput. The result is normalized for display.

Finally, the Kivy GUI displays the processed line spectrum or a rolling spectrogram. The session buffer preserves raw frames for export, and CSV output includes both raw ADC counts and derived columns such as wavelength, voltage, processed counts, and normalized intensity.

# Slide Footer / Notation

Recommended footer: lambda = wavelength; a.u. = arbitrary units.

Recommended axis labels: x-axis = pixel for uncalibrated views, lambda (nm) after wavelength mapping; y-axis = normalized intensity (a.u.) unless a physical radiometric calibration has been completed.